Corsano SDK Integration Protocol

To start with the integration, we would like to clarify requirements. The steps for a successful integration are:

1) Client informs Corsano which Vital Parameters they want to monitor and frequency

2) Corsano will analyse and propose best matching Monitoring Plan

3) Corsano prepares document confirming Monitoring Plan, Vital Parameters, Frequency

4) Both parties sign document that will be basis of the SDK integration

5) Start integration project

6) Integration with regular reporting and logs to Corsano

7) Technical validation and verification

8) Approval integration project by Corsano

9) Signing SDK License Agreement (see template)

10) Clinical implementation (cannot start before completion prior to technical approval and signing of the SDK License Agreement)

Please describe your project where you want to integrate the Corsano SDK/BLE API:

Please describe Vital Parameters and Frequency you want to monitor:

| Vital Parameter | Recommended Frequency | Client Project |
| --- | --- | --- |
| Pulse Rate | 1/5 min |  |
| SPO2 | 1/5 min |  |
| BRPM | 1/5 min |  |
| CBT | 1/5 min |  |
| Cuffless NIBP | 1/30 min |  |
| Steps | 1/5 min |  |

Please select minimal Raw Data to transfer from bracelet to app:

| Raw Data | Recommended Frequency | Client Project | Remarks |
| --- | --- | --- | --- |
| ACC | 32Hz |  |  |
| PPG G | 32Hz |  | Note 128Hz reduces battery life to 36-48h, please use only for short monitoring of patients. |
| PPG G/R/iR | 32Hz |  | Note 128Hz reduces battery life to 28-36h, please use only for short monitoring of patients. |
| BIOZ | 25Hz |  |  |

SDK Integration for:

|  | iOS |
| --- | --- |
|  | Android |

Plans description for 287-2b

| Plan ID | Plan name | Autonomy (days) | Vital parameters | Default Sampling rate of vital parameters | Raw data |
| --- | --- | --- | --- | --- | --- |
| 1 | Max battery | 15 | Steps, Calories, CBT | 1/min | - |
| 2 | Typical | 7.5 | Steps, Calories, CBT HR BRPM, RR intervals Sleep | 30s/1min/5min 30s/1min/5min 1/30min 1/s at nights 1/30s at nights | - |
| 3 | High Resolution BPM | 6 | Steps, Calories, CBT HR BRPM, RR intervals Sleep | 30s/1min/5min 30s/1min/5min 30s/1min/5min 1/s 1/30s at nights | - |
| 7 | Typical multicolor | 5.5 | Steps, Calories, CBT HR BRPM, SPO2 RR intervals Sleep | 30s/1min/5min 30s/1min/5min 30s/1min/5min 30s/1min/5min 1/s at nights 1/30s at nights | Periodic raw PPG G/R/IR 128Hz and Raw Acc 32Hz (2minutes / Every 30 minutes) |
| 8 | High Resolution Multicolor | 4 | Steps, Calories, CBT HR BRPM, SPO2 RR intervals Sleep | 30s/1min/5min 30s/1min/5min 30s/1min/5min 30s/1min/5min 1/s 1/30s at nights | Periodic raw PPG G/R/IR 128Hz and Raw Acc 32Hz (2minutes / Every 30 minutes) |
| 10 | Continuous Green, intermittent multicolor | 5 | Steps, Calories, CBT HR BRPM, SPO2 RR intervals Sleep | 30s/1min/5min 30s/1min/5min 30s/1min/5min 1/30min 1/s 1/30s at nights | Periodic raw PPG G/R/IR 128Hz and Raw Acc 32Hz (2minutes / Every 30 minutes) |
| 9 | Continuous Green, intermittent multicolor Raw | 3 | Steps, Calories, CBT HR BRPM, SPO2 RR intervals Sleep | 30s/1min/5min 30s/1min/5min 30s/1min/5min 30s/1min/5min 1/s 1/30s at nights | Continuous raw PPG G 32 Hz Periodic raw PPG G/R/IR 128Hz and Raw Acc 32Hz (2minutes / Every 30 minutes) |
| 5 | Hospital raw | 3.5 at 32Hz 1.5 at 128Hz | Steps, Calories, CBT HR BRPM, SPO2 RR intervals Sleep | 30s/1min/5min 30s/1min/5min 30s/1min/5min 30s/1min/5min 1/s 1/30s at nights | Continuous raw PPG G/R/IR , 32 Hz or 128 Hz |
| 11 | 5 min plan | 6 | Steps, Calories, CBT HR BRPM, SPO2 | 5min 5min 5min or 30min 5min or 30min |  |

Overview Raw Data

| Raw Data | Available Frequency | Recommended Frequency | Remarks |
| --- | --- | --- | --- |
| ACC | 25Hz, 32Hz | 32Hz |  |
| PPG G | 32Hz, 64Hz, 128Hz | 32Hz | Note 128Hz reduces battery life to 36-48h, please use only for short monitoring of patients. |
| PPG G/R/iR | 32Hz, 64Hz, 128Hz | 32Hz | Note 128Hz reduces battery life to 28-36h, please use only for short monitoring of patients. |
| BIOZ | 25Hz | 25Hz |  |

Acronyms:

CBT: core body temperature,

HR: heart rate, beats per minute

BRPM: breath rate per minute

RR intervals: the time interval between two consecutive heart beats

Acc: 3-axis accelerometer data.

Important Notes:

- Sampling rate: The values in the table are the default numbers.

- Raw PPG sampling rates: The raw PPG data can be set to 32Hz or to 128Hz using a ble command.

- Raw PPG channels: we have 3 channels for the PPG data:  Green, Red, infrared.

- Sleep in mainly processed in the Corsano mobile app. The bracelet only provides necessary data to the app, but not sleep data.

- Raw ACC can be recorded continously at 32 Hz in any plan as optional. It will decrease the autonomy of the bracelet by one or 2 days depending on the selected plan.

- Raw BioZ can be recorded continously at 25 Hz in any plan as optional, as while as the PPG is not recorded at 128Hz. This will decrease the autonomy of the bracelet by one or 2 days depending on the selected plan.

- Emography can be recorded continously every minute in any plan as optional, as while as the PPG is not recorded at 128Hz. This will decrease the autonomy of the bracelet by one or 2 days depending on the selected plan.